STEREO/IMPACT/SIT

SIT Door Close Procedure

Ver. 1.0  01/25/05

# Unit:______  Test:________________________  Date:________

1.  PURPOSE

This document specifies the procedure for closing the  SIT acoustic door.

Duration of procedure is about 10 minutes but it may be interleaved with other non-SIT procedures to save time.

2.  REQUIREMENTS

2.1 Equipment

1) Requires access to the SIT Actuator Connector, and an ohmmeter capable of measuring 5 ohms to +/- 20%.

2) Requires a manual TiNi reset tool

2.2  Caution

The open door exposes extremely thin and fragile Ni foils.  Take care not to let anything come in contact with the foils.  Take care not to direct a strong stream of air into the open sunshade.

2.3 Cleanliness

This procedure requires touching the acoustic door.  This must be done with xxxx gloves only.  The procedure may only be carried out in a class 10,000 or better clean environment

3. PROCEDURE

3.1	Manually close door

3.2	Insert reset tool in TiNi actuator and press with 5 lbs force to reset the actuator.

3.3	Remove tool

3.4	Measure impedance of actuator.  Expect 5ohms +/- 20%

Resistance:__________________________

End of procedure