STEREO/IMPACT/SIT

SIT HVPS In-flight

Turn-on Procedure

Date: 06/06/09

Unit: _____________________   Date: ____________________

PURPOSE

This document specifies the procedure for turning on and ramping up the SIT high voltage power supply in flight.

Assumption:  This procedure assumes that the instrument will be powered on continuously for 5 consecutive days.

I.  Perform an aliveness test before sending commands for the turn-on procedure.

Verify SIT table checksums available in housekeeping data on SIT GSE:

| Description | Expected Value | Actual Value |
| --- | --- | --- |
| SIT FM1 tab cksm | 52A82E (4/28/06) |  |
| SIT FM2 tab cksm | 46E9B7 (4/28/06) |  |

Verify the following singles rates on SIT GSE match expected values (cts/min):

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | >600 |  |  |
| VSE | 0 |  |  |
| Matrix Rates (ApId 605) | All 0 |  |  |
| Beacon Rates (ApId 619) | All 0 |  |  |
| Pulse Height Events (ApId 606-617) | All 0 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected 			values:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SW Ver | 0602 |  |  |
| SW error | 0 = (no error) |  |  |
| HV Step | 00 |  |  |
| Toferror | 0 = error events |  |  |
| HV Status | 0 = off |  |  |
| Eonly | 0 = ET coincidence required |  |  |
| Junk (Events) | 0 = junk events ignored |  |  |
| LIMHI | 500 |  |  |
| calib gain | 9 - 11 |  |  |
| calib offset | -14 to –65, should settle at –15 but may take as much as an hour (don’t wait for it!). |  |  |
| calib error | 0x08, should settle to 0x00 |  |  |

Verify the following status information has been returned on SIT GSE and matches expected values:

| Description | Expected Values | Actual Values | Comments |
| --- | --- | --- | --- |
| HV | -79.6 |  |  |
| Tof T | 25-35 C |  |  |
| Foil T | 20-25 C |  |  |
| SSD T | 20-25 C |  |  |
| +3.3 V monitor | 3.3 V |  |  |
| +2.5 V monitor | 2.5 V |  |  |
| +5 V monitor | 5 V |  |  |
| +6 V monitor | 6 V |  |  |

II.   Command sequence for HVPS turn on

**************************** Day 1 ****************************

Send the SIT Commands:

immed 1

hvenable 1

CAUTION: the following is a hazardous command.  Check carefully before sending

hvramp 40

After 2-3 minutes, verify the HV setting and count rates:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 40 |  |  |
| HV | 985 - 1045 |  |  |
| HV Status | 1 = on |  |  |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | >600 |  |  |
| VSE | 0 |  |  |

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp 80

After 2-3 minutes, verify the HV setting and count rates:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 80 |  |  |
| HV | 2063 - 2156 |  |  |
| HV Status | 1 = on |  |  |
| SRT | 0-10 |  |  |
| STP | 0-10 |  |  |
| VS | 0-10 |  |  |
| SSD | >600 |  |  |
| VSE | 0-10 |  |  |

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp 90

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | 90 |  |  |
| HV | 2345 - 2422 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >10 |  |  |
| STP | >10 |  |  |
| VS | >0 |  |  |
| SSD | >600 |  |  |
| VSE | >0 |  |  |

**************************** Day 2 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp a0

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | A0 |  |  |
| HV | 2610 - 2704 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >10 |  |  |
| STP | >10 |  |  |
| VS | >0 |  |  |
| SSD | >600 |  |  |
| VSE | >0 |  |  |

**************************** Day 3 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp a8

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | A8 |  |  |
| HV | 2760 - 2853 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

**************************** Day 4****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully.

hvramp b0

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | B0 |  |  |
| HV | 2843 - 2986 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

**************************** Day 5 ****************************

Send the SIT Command:

CAUTION: the following is a hazardous command.  Check carefully before sending

hvramp b8

After 2-3 minutes, verify the HV setting:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV Step | B8 |  |  |
| HV | 2976 - 3109 |  |  |
| HV Status | 1 = on |  |  |
| SRT | >100 |  |  |
| STP | >100 |  |  |
| VS | >1 |  |  |
| SSD | >600 |  |  |
| VSE | >1 |  |  |

******************************************************

Notes:

(1) range of expected HV values based on range between fm1 and fm2 hot/cold cases observed during thermal vac, ± 30 V allowance for 2 channel jitter in ADC

(2) range of expected count rates in detectors variable due to diferent levels of activity observed in IP space.  Figure below shows COUNTS/SECOND for Wind/STEP for comparable count rates during first portion of 2006